## Slide 1: GWAS Data Status: ATTREX 2013
E. Atlas, M. Navarro, X. Zhu, L. Pope, R. Lueb

## Slide 2: Sample Collection
GWAS capacity 90 samples / flight
RF01 Pump Failure about 1/3 into flight
RF02 Electrical failure shutdown system (no samples)
RF03 – RF06: Success (88 - 90)
Total samples: 388

## Slide 3: Field Analytical System

## Slide 5: 32 gases to be reported to archive
CCl3F_CFC_11 |  |  |  | OCS_Carbonyl_sulfide |  |  | 
CCl2F2_CFC_12 |  |  |  |  |  |  | 
C2Cl3F3_CFC_113 |  |  |  | CBrClF2_Halon_1211 |  |  | 
 |  |  |  | C2Br2F4_Halon_2402 |  |  | 
CH3CCl3_Methyl_chloroform |  |  |  |  |  |  | 
CCl4_Carbon_tetrachloride |  |  |  | CH3Cl_Methyl_chloride |  |  | 
 |  |  |  | CH3Br_Methyl_bromide |  |  | 
CHClF2_HCFC_22 |  |  |  | CH3I_Methyl_iodide |  |  | 
CH3CClF2_HCFC_142b |  |  |  |  |  |  | 
CH3CCl2F_HCFC_141b |  |  |  | C2H6_Ethane |  |  | 
 |  |  |  | C2H2_Ethyne |  |  | 
CH2Cl2_Methylene_chloride |  |  |  | C3H8_Propane |  |  | 
CHCl3_Chloroform |  |  |  | C6H6_Benzene |  |  | 
C2H4Cl2_1_2_Dichloroethane |  |  |  |  |  |  | 
C2Cl4_Tetrachloroethylene |  |  |  | CH2BrCl_Bromochloromethane |  |  | 
C2HCl3_Trichloroethylene |  |  |  | CH2Br2_Methylene_bromide |  |  | 
C6H5Cl_Chlorobenzene |  |  |  | CHBrCl2_Bromodichloromethane |  |  | 
 |  |  |  | CHBr2Cl_Dibromochloromethane |  |  | 
CH3ONO2_Methyl_nitrate |  |  |  | CHBr3_Bromoform |  |  | 
C2H5ONO2_Ethyl_nitrate |  |  |  |  |  |  | 
C3H7NO3_Isopropyl nitrate |  |  |  |  |  |  | 

## Slide 6: Analytical Issues
Artifact in ethyne: coelution of unidentified CFC/HCFC
  Current version subtracts constant amount (22 pptv)
ATTREX (2013)
+ ATTREX (2011)

## Slide 7: Analytical Issues (2)
Moisture related artifact during sample injection
  Sample response different from standard response
  Use long-lived tracer to normalize responses
    CFC114 (lifetime ~ 189 yr)
    Expected variation ~1%/100 – 150 ppb O3
  Results in some loss of samples when CFC114 analytical problem (H2O)
  Effect depends on volatility (GC retention)
    less effect for more volatile than 114
      CFC12, CH3Cl, HFC134a
  Introduces extra imprecision into long-lived tracers; less of an impact for short-lived substances

## Slide 8: Example (RF05) : Methyl nitrate

## Slide 9: Example (RF05) : Methyl nitrate

## Slide 10: Analytical Issues (3)
Original working standard vented during shipment.
  Last minute replacement needed calibration
  Still some standard checks required with primary standards when SEAC4RS analysis complete.

## Slide 11: Data to archive after meeting